## SEM-C02
S-9220 CONFIGURATION
1-Supplier
Name | HITACHI
2-Equipement Type
Description | S-9220
3-Process
Process type | CD-SEM
4-Wafer Type
Wafer specification | 200mm (8in) Notch
5-Layout configuration
Specific Layout defintion | Standard
6-CE Compliance
7-Mainframe configuration
Platform
configuration
Position 1 | Body
(1) Comprehensive specifications
・Wafer size | Wafer diameter 200 mm
・Length measurement method
・Measuring range | 0.1μm～2.0μm
・Length measurement reproducibility | ± 1% or 3nm (3σ) larger value
・Throughput | Continuous measurement mode 56 wafers / hour (using our standard wafers,
our measurement conditions)
・Secondary electron image resolution | 3 nm (acceleration voltage 0.8 kV)
・Image magnification | Optical microscope image: Approximately 110 times
SEM image : 1k to 300k times
・Length measurement reproduction accuracy | 40k-150k-200k times (However, for magnifications exceeding 150k times,
guaranteed magnification | The effect of electron beam irradiation shall be excluded. )
・Measurement mode | Automatic measurement, manual measurement
(2) Stage
・Range of movement | X-axis: 0-200 mm Y-axis: 0-200 mm
・Drive system | Pulse motor drive
・Control | Computer control Positioning control by linear encoder
・Drive speed | Maximum 100 mm / sec (each X-axis and Y-axis)
(3) Loader system
・Wafer carrier → loader room | Automatic transfer by wafer transfer robot
・Between loader room and stage | Automatic vacuum exhaust, automatic loading
(4) Wafer transfer robot
・Method | Random access type for 2 cassettes
・Wafer detection in cassette | Wafer searcher built-in
・Chucking method | Vacuum suction on the back side of the wafer
・OF, V-notch detection | Automatic detection by optical sensor
(5) Electron optics
・Acceleration voltage | 500 to 1600V (variable in 10V steps)
・Lens system | Electromagnetic lens system, booster stock lens
・Detection system | Scintillator / photomultiplier detector using EXB filter.
Two-stage detection method for SE and BSE images
・Objective lens diaphragm | Heated thin film movable pattern (4 holes switching from outside the vacuum)
with fine adjustment mechanism
・Scanning coil | 2-stage electromagnetic deflection
・Astigmatism correction | 8-pole electromagnetic system (X, Y)
・Prove current monitor | Built-in Faraday cup, with automatic measurement function
・Optical microscope | Approximately 1.2 mm square field of view, with black and white CCD camera
(6) Control / display system
・Observation and control CRT | EWS (GUI) 21-inch CRT
Combined use of SEM image / optical image display and computer dialogue operation
Display measurement conditions, wafer map, length measurement value,
stage coordinates, etc.
・Scanning mode | TV scan, with auto brightness / contrast function
・Image processing | Image processing by software using filtering (standard equipment).
Real-time image processing (optional) using DSP.
・Image recording | Video printing device (optional),
Image file save function (optional)
・Safety device | With emergency shutoff switch (mushroom type)
(7) Length measurement data processing
・ File save function | Various measurement condition file save function
・ Storage medium | 9G byte hard disk (including for system programs)
・ Auxiliary storage medium | 3.5-inch magneto-optical disk driver, 3.5-inch F / D driver
・Data embedding function | Statistical calculation function using worksheets (mean value, standard deviation, etc.)
Real-time length measurement value graph display function
・Print output | 80-digit thermal printer
(8) Vacuum exhaust system
・Method | Fully automatic dry clean exhaust system
・Vacuum pump | 3 ion pumps Two turbo molecular pumps 2 dry pumps (supplied)
・Safety function | Protection function in the event of a power outage, AIR pressure drop,
and cooling water flow rate drop (self)
With dynamic valve closing function, pump stop function)
9-Chemical Solution configuration
10- Workstation hardware configuration
11-System software/ automation configuration
12-Electrical
13-Ancillaries
14- Hook-up connexion / Specific requirement
15-Options
17- Target process performances

## xl_DCF_History (hidden)
CLINAME | DATETIME | DONEBY | IPADDRESS | APPVER | RANDOM | CHECKSUM
ᕧᖀᕕᕾᕳᖅᖅᕻᕸᕻᕷᕶ | ᕈᕁᕃᕆᕁᕄᕂᕃᕊᔲᔲᕃᕂᕌᕂᕉᕓᕟᔲᔺᕙᕟᕦᔽᕄᕌᕂᔻ | ᕥᕦᕮᕔᕷᖀᖁᕻᖆᔲᕿᕳᖀᕻᖀᕻ | ᕤᕡᕧᕃᕄᕉᕄᕄ | ᕈᕀᕂᕀᕂᕀᕂ | 18 | ᕆᕉᕂᕈ
౞౷ౌ౵౪౼౼౲౯౲౮౭ | ఼సాస఻హ఻఺఩఩఺ిృ఼఼఩ఱ౐ౖౝఴ఺ృహల | ౤౞౷ౌ౵౪౼౼౲౯౲౮౭౦఩ే఩ొ౒ౙ఩౱౪౷౭౸౿౮౻ | ౌ౛ౘౌౠౕ఺఼ూహూ | ీష఺షహషహ | 9 | ాు఻హ