| 参数要求：Parameter requirements: |  |
| --- | --- |
| 温度波动度 Temperature fluctuation degree | ±0.5℃（空载） ±0.5℃(No load) |
| 温度均匀度 Uniformity of temperature | ±2℃（空载） ±2℃(No load) |
| 湿度波动度 Humidity fluctuation degree | ±2℃（空载） ±2℃(No load) |
| 湿度均匀度 Uniformity of humidity | 3-4%（空载） 3-4%(No load) |
| 升温速率 Ramp rate | 1℃~5℃/min |
| 降温速率 cooling rate | 0.7~4℃/min |
| 仪表精度 Instrument accuracy | 0.01℃ |
| 噪音 Noise | ≤50dB |
| 功率 power | 3~10KW |
| 电源 Power supply | 220V/380V 50Hz |
| 箱体规格：Box dimensions |  |
| 外壳 shell | 1.5mm厚冷轧钢板 1.5mm thick cold-rolled steel plate |
| 颜色 Color | 乳白色环保烤漆 Milk white eco-friendly baking varnish |
| 内胆 inner liner | 1.2mmSUS304不锈钢拉丝板 1.2mm SUS304 stainless steel brushed plate |
| 保温层 Insulation layer | 100mm超细保温玻璃棉，高温低温箱体外部无发烫和凝露 100mm superfine thermal insulation glass wool, no scalding or condensation on the exterior of the high and low temperature chamber. |
| 箱体隔层 Box compartment partition | 两层不锈钢网状结构，可随意调节高度。 The two-layer stainless steel mesh structure can be adjusted in height at will. |
| 滑轮pulley | 全方位定位轮All-round positioning wheel |
| 玻璃窗glass window | 钢化导热玻璃窗Tempered heat-conducting glass window |
| 铰链 Hinge | 不锈钢材质 Stainless steel material |
| 照明Illumination | 节能照明，能清楚看到箱体内试验产品情况 Energy-saving lighting allows for a clear view of the test products inside the box. |
| 硅胶条 Silicone strip | 抗老化高温硅胶条可耐400℃高温和-180℃低温 Anti-aging high-temperature silicone strips can withstand temperatures as high as 400℃ and as low as -180℃. |
| 加湿器 Humidifier | 内置不锈钢加湿器 Built-in stainless steel humidifier |
| 加湿水箱 Humidification water tank | 大容量不锈钢水箱 Large-capacity stainless steel water tank |
| 控制配置：Control configuration: |  |
| 控制仪表 Control instrument | 台湾“威纶”可编程智能仪表，彩色全触摸屏控制，中英文切换，定值运行和程式运行均可。可设定曲线不少于30条程序曲线，每条曲线最大设定不少于800步，一次可连接不少于6个程序，最多可做不少于10个的连接，单一程序可做不少于999次的重复试验。可实时显示程序的运行时间、段数、剩余时间、重复次数：可显示试验数据，包括设定温度、实测温度、总运行时间、段运行时间、段剩余时间、加热状态、日历时间等；温度直接数字显示等。并可配有数据接口。 Taiwan's "Weineng" programmable intelligent instrument features a full-color touch screen control with the ability to switch between Chinese and English. It supports both fixed value operation and program operation. It can set no less than 30 program curves, with each curve capable of having no less than 800 steps. It can connect no less than 6 programs at a time and make no less than 10 connections in total. A single program can be repeated no less than 999 times. It can display the running time, segment number, remaining time, and repetition count of the program in real time. It can also show test data, including set temperature, measured temperature, total running time, segment running time, remaining time of the segment, heating status, calendar time, etc. The temperature is directly displayed in digital form. It can be equipped with a data interface. |
| 固态继电器 Solid state relay | 施耐德、德力西、台湾扬明等知名品牌 Well-known brands such as Schneider, Delixi, and Taiwan Yangming |
| 交流接触器 AC contactor | 施耐德等 Schneider et al. |
| 其他原件 Other original documents | 均为知名品牌 All are well-known brands. |
| 制冷配置：Refrigeration configuration: |  |
| 机组 crew | 1) 采用法国“泰康”全封闭制冷压缩机组，每台机组均经过欧洲“泰康”电脑联网逐项监测并有防伪编码，可通过电脑上网查寻 1) It adopts the fully-enclosed refrigeration compressor units of "Tecumseh" from France. Each unit has been monitored item by item through the "Tecumseh" computer network in Europe and has an anti-counterfeiting code, which can be searched online via the computer. |
|  | 2) 制冷压缩机：为保证试验箱对降温速率和可达最低温度的要求，该试验箱采用单组式/复叠式制冷系统，复叠式系统包含一个高压制冷循环和一个低压制冷循环，其连接容器为蒸发冷凝器，蒸发冷凝器的功能为将低压循环的蒸发器作为高压循环的冷凝器之用 2) Refrigeration compressor: To ensure that the test chamber meets the requirements for cooling rate and the lowest achievable temperature, this test chamber adopts a single-stage or cascade refrigeration system. The cascade system consists of a high-pressure refrigeration cycle and a low-pressure refrigeration cycle, with an evaporative condenser as the connecting container. The function of the evaporative condenser is to use the evaporator of the low-pressure cycle as the condenser of the high-pressure cycle. |
| 蒸发器 Evaporator | 专业定制 Professional customization |
| 制冷剂 Refrigerant | 环保404A低压循环 R23高压循环 Environmental protection 404A low-pressure cycle R23 high-pressure cycle |
| 安全保障：Safety guarantee: |  |
| 1) 超温报警 Over-temperature alarm |  |
| 2) 漏电保护 Leakage protection |  |
| 3) 欠相缺相保护 Under-voltage and phase loss protection |  |
| 4) 过电流保护 Overcurrent protection |  |
| 5) 快速熔断器 Fast-acting fuse |  |
| 6) 缺水保护 Water shortage protection |  |
| 7) 压缩机高低压保护 High and low pressure protection for compressors |  |
| 8) 压缩机过热保护 Compressor overheat protection |  |
| 9) 压缩机过电流保护 Compressor overcurrent protection |  |
| 10) 线路保险丝及全护套式端子Line fuses and fully insulated terminals |  |
| 11) 接地保护 Grounding protection |  |