- This device adopts energy dispersive X-ray fluorescence analysis method, which can analyze a wide range of elements with low detection limits and is suitable for various sample shapes and atmosphere conditions. It is equipped with a high-purity silicon detector and a liquid nitrogen cooling system to ensure efficient and accurate detection. It has an optical image observation function and the software supports qualitative and quantitative analysis. The system requires a high-configuration computer and specific environmental conditions, and is suitable for laboratory use.

| 分析原理 Analysis principle |  | 能量色散X射线荧光分析法 Energy dispersive X-ray fluorescence analysis method |
| --- | --- | --- |
| 分析元素 Analysis of elements |  | Si～U(Cd/Pb高精度型)High-precision type |
|  |  | Na～U(φ1.2 mm/φ3 mm切换方式型) Switching mode type |
| 检出下限(Cd/Pd) Detection limit (Cd/Pd) |  | Cd≦2 ppm, Pb≦5 ppm Cl≦50ppm |
| 样品形状 Sample shape |  | 最大460×360 mm(高150 mm) maximum 460×360 mm(hight 150 mm) |
| 样品室气氛 Atmosphere in the sample room |  | 大气 Atmosphere |
| X射线管 X-ray tube | 靶材 Target material | Rh |
|  | 管电压 Tube voltage | 最大50KV maximum 50KV |
|  | 管电流 Tube current | 最大1 mA maximum 1 mA |
| X射线照射径 X-ray irradiation path |  | Φ3 mm (同时搭载：φ1.2mm ) Φ3 mm equipped with simultaneously φ1.2mm |
| 检测器 detector |  | 高纯硅检测器XEROPHY High-purity silicon detector XEROPHY |
| 液氮罐容量 Liquid nitrogen tank capacity |  | 3升 3L |
| 液氮消耗量 Liquid nitrogen consumption |  | 每日1升以下（不开机时不用） Less than 1 liter per day (not in use when the machine is off) |
| 光学图像观察 Optical image observation |  | 倍率50倍（X射线同轴） Magnification: 50 times (X-ray coaxial) |
| 软件 Software |  | 定性分析：自动定性或手动定性 Qualitative analysis: automatic qualitative or manual qualitative |
|  |  | 定量分析：检量线法 Quantitative analysis: calibration curve method |
|  |  | 照射直径3mm时 When the irradiation diameter is 3mm |
|  |  | 基础参数法（1）标准法 (2)1点校正 Basic parameter method (1) Standard method (2) One-point correction |
|  |  | 用户设定功能：可以按照用户事先编好的条件，按照程序自动进行分析和定量分析 User-defined function: It can automatically conduct analysis and quantitative analysis based on the conditions pre-set by the user. |
| 计算机 computer | CPU | Pentium IV 1.8 GHz以上 Pentium IV 1.8 GHz or above |
|  | 内存 | 1GB以上 More than 1GB |
|  | 硬盘 | 120GB以上 Over 120GB |
|  | OS | Windows XP |
| 监视器 Monitor |  | 17寸 CRT（任选：17寸 LED） 17-inch CRT (optional: 17-inch LED) |
| 打印机 printer |  | 彩色喷墨打印机 Color inkjet printer |
| 周围温度 ambient temperature |  | 10～35 ℃(性能温度)/ 5～40 ℃(动作温度) 10 to 35 ℃ (operating temperature) / 5 to 40 ℃ (action temperature) |
| 周围湿度 Surrounding humidity |  | 5～31 ℃时温度范围：最大相対湿度80 %以下 Temperature range: 5 - 31 ℃; Maximum relative humidity: below 80% |
|  |  | 31～40 ℃时温度范围：相対湿度50 %以下 Temperature range: 31 - 40 ℃; Relative humidity: below 50% |
| 电源 Power supply |  | AC 100 V、120 V、220 V、240 V ±10 %、50/60 Hz |
| 消耗电力 consume electricity |  | 1.3 kVA 以下(含计算机、CRT、打印机) Below 1.3 kVA (including computers, CRTs, printers) |
| 设备重量 Equipment weight |  | 约 265 kg(不含桌子、计算机) Approximately 265 kg (excluding the table and computer) |
| 外形尺寸 Overall dimensions | 分析单元Analysis unit | 610(W)×750(D)×500(H)mm |
|  | 信号处理单元 Signal processing unit | 220(W)×500(D)×480(H)mm |