## Pre-Check
System Health Pre-Check
Software release | 8.9.0
Machine pending issue
Customer support engineer
Patch level | 97c
Parts status
Tool status
Item | Value | Pre Check
Slit Uniformity | Annular:0.8/0.85/0.55 | 0.55
Intensity | Annular | 1643
Dose Accuracy | Repeatability | 0.09
| Accuracy | -0.36
Wafer Level Telecentricity | Rx/Ry | 4.07
REMA Offset | Calibrated offset X0
Calibrated offset X1
Calibrated offset Y0
Calibrated offset Y1
TIS Repro | Max3 * STD pos.x [nm] | 7.3, 6.1, 7.1, 6.8, 6.6
| Max3 * STD pos.y [nm] | 4.5, 3.5, 2.5, 3.9, 4
| Max3 * STD pos.zx [nm] | 22.4, 20, 21.9, 20.1, 20.5
| Max3 * STD pos Zy [nm] | 15.8, 18.1, 17, 15.8, 17
| Failed scans | 0,0,0,0,0
Corse Skew | Reticle to RS rotation Std dev | 0.4
| Reticle to WS rotation Std dev | 1
| Beam Length at zero Rz_x - New | 11.4
| Beam Length at zero Rz_y - New | 11.4
| Mirror to reticle rotation | 15.7
| RS to WS rotation (skew) | 21.3
| Reticle to RS rotation | -3.9
SWS Horizontal Servo Perfromance | MA-X | 4.3
| MA-Y | 4.7
| MA-Rz | 29.7
| MSD-X | 22
| MDS-Y | 36
| MSD-Rz | 90
MASR | Xt | 2.61
12.79
| Yt | 2.61
9.57
| Zr | 2.79
16.54
| Zt | 1.92
8.5
| Yr | 1.05
2.44
| Xr | 0.87
2.44
SWDP | MA-X | 1.1
| MA-Y | 0.6
| MA-Rz | 19.2
| MSD-X | 10.2
| MSD-Y | 9.9
| MSD-Rz | 171.3
SRUC | Max. Peak Position Err- all pos.
| X (nm) | 33.8
| Y (nm) | 15.4
| Rz (nm) | 144
| Z (um) | 0.064
| Rx (urad) | 0.576
| Ry (urad) | 0.589
Max. accuracy and reproducibility During Scan - Mean
| MA-X (nm) | 3.5
| MA-Y (nm) | 1.3
| MA - Rz (nrad) | 10.8
| MA - Total Z (nm) | 32.8
| MSD - X (nm) | 15.8
| MSD-Y (nm) | 6.9
| MSD - Rz (nrad) | 58.7
| MSD - Total Z (nm) | 66.8
Measure the stray light with SAMOS | Cleaning trigger max stray light at reference blocksize [%] | 1.12
Straylight with Rema open and close | Maximum Stray Light Number [%] | 0.53
TIS pupil | Graphic Ok or not ?
| Sigma in | 0.139
| Sigma out | 0.704
| Ellipticity | -81.9%
| X/Y balance | .152% and 7.175%
OA wafer quality | Max NF-r | 3057
| Max NF-g | 2079
OA repro | SD X and Y values | yes
Zero calibration | Home to Zero distance X [mm] | -15.044
| Home to Zero distance Y [mm] | 30.689
| Home to Zero_angle [mrad] | 0.509
| Zero sensor slope X [nm/bit] | 17.244
| Zero sensor slope Y1 [nm/bit] | 16.198
| Zero sensor slope Y2 [nm/bit] | 16.905
| Zero to Init Angle [mrad] | 0.509
Initial Vertical position:
| Z[mm] | 0
| Rx[mrad] | 0
| Ry[mrad] | 0
| Y1-Y2 Zero Sensor Distance[mm] | 175
Takeover qualification | Measurement reproducibility:
| x SDEV | 0.1
| y SDEV | 0
| phi SDEV | 2
Swap up to takeover to dipod:
| dx SDEV | 0.4
| dy SDEV | 0.4
| dphi SDEV | 8.1
Swap up to dipod rotation:
| dx SDEV | 0.6
| dy SDEV | 1.1
| dphi SDEV | 20.8
Swap up to E-pins:
| dx SDEV | 0.9
| dy SDEV | 0.7
| dphi SDEV | 40
Swap up to E-table:
| dx SDEV | 0.9
| dy SDEV | 1.1
| dphi SDEV | 24
Edge sensor repro | Standard dev. Xalign | 1.66
| Standard dev. Yalign | 0.88
| Standard dev. rot | 29.72
LS height and tilt | Screw values A, | 8
| Screw values B | 4
| Screw values C | 7
LS performance
Measurement [V]:
Reference [V]:
Measurement [V]:
| Measurement [V]: | -7.482
| Measurement [V]: | 2.402
Check Focus Reproducibility | Absolute Std. Dev. Total Z [um] | 0.0108
| Absolute Std. Dev. Rx [urad] (along slit axis) | 0.271
| Absolute Std. Dev. Ry [urad] (along scan axis) | 0.322
Image quality reproducibility | Magnification [ppm] | 0.3
| Rotation [urad] | 0.17
| Translation X [nm] | 3.66
| Translation Y [nm] | 2.49
| Image Tilt Rx [urad] | 3.63
| Image Tilt Ry [urad] | 0.83
| Focus [nm] | 11.93
| StoLe Tilt Rx [urad] | 0.39
| StoLe Tilt Ry [urad] | 0.3
| Actuator Tilt Rx [urad] | 0.28
| Actuator Tilt Ry [urad] | 0.27
Global align repro(OA) | Scan scaling [ppm] | 0.04
| Scan skew [μRad] | 0.02
| Reticle translation X | -0.77
| Reticle translation Y | -2.28
| Reticle mag | 0.039
| Reticle rotation | 0.033
| Wafer translation X | 0.005
| Wafer translationY | 0.006
| Wafer magnification | 0.029
| Wafer rotation | 0.023
Image Plane Deviation / AST | Image Plan Deviation | 11
| Astigmatism | 5
| Image Distortion Non Correctable Error | X=6.6/Y=6.9
Stage accuracy | Stage Repeatability X and Y | 2.7
| Stage Repeatability Y and Y | 2.1
stage grid verification | Vector, 3 sigma;XX | 13.3
| Vector, 3 sigma;XY | 6.5
| Vector, 3 sigma;YX | 21
| Vector, 3 sigma;YY | 5.9